## Slide 1
David Kessler – University of Massachusetts, Amherst
Muon g-2 Collaboration
CIPANP 2025
11 June 2025
Precision Magnetic Field Measurements in Muon g-2

## Slide 2: The g-2 Magnetic Field
The Muon g-2 experiment uses a magnetic storage ring at 1.45 Tesla.
Muon anomalous moment, equation for :
precision goal: 140 ppb.
precision goal: 70 ppb.
Muon g-2 Magnetic Storage Ring.

## Slide 3: Measuring with Nuclear Magnetic Resonance (NMR)
FID frequency is proportional to B.
FID signals from proton precession in sample.
g-2 Calibration Probe Diagram.
Coefficient known to 4.1 ppb.

## Slide 4: Measurement Overview
Trolley Probes and Fixed Probes:
17 trolley probes map the field along the muon beam path.
378+ fixed probes track changes in the field while muons are running.
Relative positions of trolley and fixed probes.

## Slide 5: Trolley Runs
Trolley readings per slice are fit to 2D multipole expansions.
Each trolley run covers the circumference of the storage ring.
Trolley runs were performed every ~3 days, with muon beam off.

## Slide 6: Analysis Overview
Tasks and Challenges:
  Interpolation – Combining trolley and FP data.
  Calibration – Trolley perturbs field while present.
  Transient Fields – Synched to muon injection, while trolley/FP are asynchronous.
The Objective:
Shielded proton
Spherical water sample at 25° Celsius
In field averaged over time/space muon distribution

## Slide 7: Using
Ratio from our collaboration’s analysis.
Precisely-known ratios from other experiments (CODATA).
Shielded proton frequency combines with , alongside known fundamental constants, to produce .
(Knowing this ratio at 25° C is why we aim for .)

## Slide 8: Interpolation
Fixed probes measure between each trolley run pair, showing evolution from one map to the next.
Any mismatches between evolved initial map and next map become uncertainties.

## Slide 9: Trolley Calibration
Determines field perturbation from trolley structure.
Each trolley probe compared to calibration probe.
  Cali probe is solo, so less field perturbation.
Alignment and Rapid Swapping procedure.
Nine campaigns over the course of g-2 operations.

## Slide 10: Trolley Calibration Campaign Results

## Slide 11: Absolute Calibration
Getting (shielded spherical-water…) from .
Tests performed with Argonne magnet to quantify each effect: Calibration Probe vs Absolute Probe.
Material Effects
~10 ppb
Shape and Susceptibility Effects
~1500 ppb
Temperature Correction
~10 ppb

## Slide 12: Cross-Calibrations
Extra tests using probes undergoing independent calibrations.
Cross-Calibration #1: Cali probe vs 3He probe.
Cross-Calibration #2: Cali probe vs J-PARC probe.
Increased uncertainty by 25ppb, but improved our confidence in the resulting uncertainty.
3He-based spherical alternate probe
Agreement on the 1.7 -level, supporting results for material effect and shape factor.
J-PARC Continuous Wave NMR probe
2019 (1.45T), 2019 (1.7T) campaigns: some inconsistencies.
2022 (3T), 2023 (1.45T) campaigns: in good agreement.

## Slide 13: Transient Fields
Kicker transient field and Faraday Effect magnetometers.
Electrostatic Quadrupoles transient field and NMR magnetometer.

## Slide 14: Uncertainty Improvement
Electric Field
Pitch
Phase Acceptance
Differential Decay
Muon Loss
Transient Kicker
Transient ESQ
Run-1 Uncertainties Table
Run-4/5/6 Uncertainties Table
Field uncertainty total: 56 ppb.
Target Achieved!

## Slide 15: Relative Field Uncertainties

## Slide 16: Conclusion
Muon g-2 result diagram:
Magnetic field measurements are important to Muon g-2’s precision goal.
NMR systems: Fixed Probes tracking, Trolley Probes mapping, and Calibration Probes calibrating.
  Plus plenty of analysis, extra measurements, and corrections!
Goal achieved!